## Instructions
Torque wrench calibration record
How to use this record
1. Fill in the yellow cells at the top: the wrench, its lowest marked value and maximum torque, the tolerance, the analyzer you are using and the date.
2. Tolerance: ISO 6789-1:2017 allows ±4% for Type II Class A, B and C wrenches with a maximum above 10 N·m, and ±6% at or below 10 N·m. Classes D to G allow ±6%. Use the maker's figure if it is tighter.
3. Set the wrench to each test point: the lowest marked value, then 60% and 100% of maximum. The set values fill in from the range you entered.
4. Take five readings at each point on a torque analyzer and type them into the As found table before you adjust anything.
5. Each reading is checked on its own. One reading outside the tolerance fails the wrench.
6. If it fails and you adjust it, repeat all three points in the As left table. The status uses the As left result when there is one.
Good practice
Use an analyzer whose error is no more than a quarter of the wrench tolerance (±1% for a ±4% wrench), with a current calibration certificate. Without your own procedure, ISO 6789-2 allows up to 24 months between analyzer calibrations.
Without your own procedure, ISO 6789-2:2017 clause 4.1 gives 12 months or 5,000 cycles, whichever comes first, counted from first use. Check sooner after an overload above 125% of maximum, a repair, a drop or other mishandling.
Store click wrenches at the lowest marked setting, never below it, and never use one to break a fastener loose.
This is an in-house check against the ISO 6789-1 limits. It is not an ISO 6789-2 calibration certificate, which also states the measurement uncertainty and comes from a calibration lab.
Track every torque wrench's due date and keep this record with it, free for up to 50 instruments: axiospec.com/signup
More free calibration tools: axiospec.com/tools

## Record
Torque wrench calibration record
In-house check of a click wrench against the ISO 6789-1:2017 limits. Yellow cells are inputs.
Results
Wrench ID |  |  | As found
Make and model |  |  | As left
Serial number |  |  | Status
ISO 6789 type and class |  |  | Largest deviation as found (%)
Lowest marked value |  |  | Largest deviation as left (%)
Maximum torque
Units | N·m
Drive size
Direction checked
Tolerance (± %) | 4
Torque analyzer ID
Analyzer calibration due
Checked by
Date checked
Interval (months) | 12
Next due
As found: readings before any adjustment
Test point | Set value | Reading | Analyzer reading | Deviation (%) | Result | Size of deviation (%)
Lowest marked value |  | 1
Lowest marked value |  | 2
Lowest marked value |  | 3
Lowest marked value |  | 4
Lowest marked value |  | 5
60% of maximum |  | 1
60% of maximum |  | 2
60% of maximum |  | 3
60% of maximum |  | 4
60% of maximum |  | 5
100% of maximum |  | 1
100% of maximum |  | 2
100% of maximum |  | 3
100% of maximum |  | 4
100% of maximum |  | 5
As left: fill in only if the wrench was adjusted
Test point | Set value | Reading | Analyzer reading | Deviation (%) | Result | Size of deviation (%)
Lowest marked value |  | 1
Lowest marked value |  | 2
Lowest marked value |  | 3
Lowest marked value |  | 4
Lowest marked value |  | 5
60% of maximum |  | 1
60% of maximum |  | 2
60% of maximum |  | 3
60% of maximum |  | 4
60% of maximum |  | 5
100% of maximum |  | 1
100% of maximum |  | 2
100% of maximum |  | 3
100% of maximum |  | 4
100% of maximum |  | 5
Deviation (%) = (analyzer reading - set value) / set value x 100, as in ISO 6789-1. Each reading must be inside the tolerance on its own.

## Example
Torque wrench calibration record
In-house check of a click wrench against the ISO 6789-1:2017 limits. Yellow cells are inputs.
Results
Wrench ID | TW-014 |  | As found
Make and model | Click wrench, 3/8 in drive (example) |  | As left
Serial number | EX-20481 |  | Status
ISO 6789 type and class | Type II Class A |  | Largest deviation as found (%)
Lowest marked value | 20 |  | Largest deviation as left (%)
Maximum torque | 100
Units | N·m
Drive size | 3/8 in square
Direction checked | Clockwise
Tolerance (± %) | 4
Torque analyzer ID | TA-02
Analyzer calibration due | 2027-03-15
Checked by | J. Rivera
Date checked | 2026-10-31
Interval (months) | 12
Next due
As found: readings before any adjustment
Test point | Set value | Reading | Analyzer reading | Deviation (%) | Result | Size of deviation (%)
Lowest marked value |  | 1 | 20.46
Lowest marked value |  | 2 | 20.38
Lowest marked value |  | 3 | 20.61
Lowest marked value |  | 4 | 20.52
Lowest marked value |  | 5 | 20.44
60% of maximum |  | 1 | 61.32
60% of maximum |  | 2 | 61.05
60% of maximum |  | 3 | 61.58
60% of maximum |  | 4 | 61.21
60% of maximum |  | 5 | 61.44
100% of maximum |  | 1 | 103.4
100% of maximum |  | 2 | 104.1
100% of maximum |  | 3 | 103.7
100% of maximum |  | 4 | 103.6
100% of maximum |  | 5 | 103.9
As left: fill in only if the wrench was adjusted
Test point | Set value | Reading | Analyzer reading | Deviation (%) | Result | Size of deviation (%)
Lowest marked value |  | 1 | 20.18
Lowest marked value |  | 2 | 20.25
Lowest marked value |  | 3 | 20.12
Lowest marked value |  | 4 | 20.3
Lowest marked value |  | 5 | 20.21
60% of maximum |  | 1 | 60.42
60% of maximum |  | 2 | 60.65
60% of maximum |  | 3 | 60.38
60% of maximum |  | 4 | 60.57
60% of maximum |  | 5 | 60.49
100% of maximum |  | 1 | 100.8
100% of maximum |  | 2 | 101.3
100% of maximum |  | 3 | 100.5
100% of maximum |  | 4 | 101.1
100% of maximum |  | 5 | 100.9
Deviation (%) = (analyzer reading - set value) / set value x 100, as in ISO 6789-1. Each reading must be inside the tolerance on its own.
Example data, made up for illustration. Reading 2 at 100 N·m is 104.10 N·m, which is +4.10% and outside ±4%, so the wrench failed as found. It was adjusted and passed as left.