## Instructions
Measurement uncertainty budget template (GUM)
This workbook does the arithmetic of a GUM uncertainty budget (JCGM 100:2008). It has no macros and works offline.
Sheets
Budget: a blank budget with 20 contributor rows. Fill the yellow cells. Gray cells are calculated.
Example (filled): an illustrative budget for a 0-25 mm outside micrometer checked against a gauge block.
Its figures show how the sheet works. They are not a budget for any real instrument.
Filling in a row
Contributor: the source of the uncertainty, for example the reference standard, the resolution or the repeatability.
Type: A when you evaluated it from repeated readings, B for any other source. It is for your record only.
Quoted value: the figure as you have it, as a number above zero. For example a certificate U, a ± half-width or an s.
Distribution: pick it from the list. It sets the divisor that turns the quoted value into a standard uncertainty.
Normal k=2: an expanded uncertainty quoted at k=2, such as a certificate U. Divisor 2.
Normal k=1: one standard deviation. Divisor 1.
Rectangular: a ± half-width where any value inside is equally likely. Divisor √3.
Triangular: a ± half-width where values near the center are more likely. Divisor √6.
U-shaped: a ± half-width where values near the limits are more likely, such as a cycling temperature. Divisor √2.
Standard: the value is already a standard uncertainty. Divisor 1.
Sensitivity coefficient c: converts the contributor into the units of the measurand. Leave it at 1 when it is already
in those units. The sheet uses the absolute value ¦c¦. A blank c counts as 1.
Degrees of freedom: optional. For a Type A standard deviation from n readings, enter n - 1. Leave it blank for infinite.
Any entry must be a number of 1 or more. The Check column flags anything else, such as a 0 or text.
A row counts once it has a quoted value above zero and a distribution. Empty rows stay blank.
What the sheet calculates
u_i = quoted value / divisor × ¦c¦, in the units of the measurand.
Combined standard uncertainty u_c = √(Σ u_i²). This root-sum-square treats the contributors as uncorrelated.
% contribution = u_i² / Σ u_i², the share of each row in the combined variance.
Expanded uncertainty U = k × u_c.
Coverage factor k
Fixed k is the default. Enter k in the cell below the coverage mode. It starts at 2.
k = 2 gives about 95% coverage when the combined distribution is approximately normal.
From effective degrees of freedom (95%): the sheet computes v_eff with the Welch-Satterthwaite formula,
v_eff = u_c⁴ / Σ(u_i⁴ / v_i). A row with no degrees of freedom counts as infinite and adds nothing to the sum.
k is Student's t for 95%, two-sided, at v_eff truncated to the next lower integer. GUM G.4.1 allows this.
When no row has degrees of freedom, v_eff is infinite and k is the normal value 1.96.
While the Check column flags a row, v_eff, k and U show a message instead of a number. Fix or clear the entry.
With Fixed k, k and U still compute. Only v_eff shows the message.
The Axiospec Welch-Satterthwaite calculator interpolates t for a fractional v_eff instead, so its k can differ slightly.
Limits
The workbook does the arithmetic. Choosing the contributors, distributions, sensitivity coefficients and degrees
of freedom is your judgement.
Correlations between contributors are not modelled.
It does not validate a budget for accreditation.
The online calculator at axiospec.com/tools/measurement-uncertainty-budget-calculator uses the same divisors and combination.
Free template by Axiospec, calibration management software: axiospec.com

## Budget
Measurement uncertainty budget (GUM)
Fill the yellow cells. Gray cells are calculated. A row counts once it has a quoted value and a distribution.
Instrument / measurement
Measurand
Units | [6 empty] | Units of the measurand. Each u_i is in these units once c is applied.
Prepared by
Date
# | Contributor (source) | Type (A/B) | Quoted value | Distribution | Divisor | Sensitivity coefficient c | Standard uncertainty u_i | Degrees of freedom v_i (blank = infinite) | u_i² | % contribution | u_i⁴ / v_i | Check
1 | [5 empty] | 1
2 | [5 empty] | 1
3 | [5 empty] | 1
4 | [5 empty] | 1
5 | [5 empty] | 1
6 | [5 empty] | 1
7 | [5 empty] | 1
8 | [5 empty] | 1
9 | [5 empty] | 1
10 | [5 empty] | 1
11 | [5 empty] | 1
12 | [5 empty] | 1
13 | [5 empty] | 1
14 | [5 empty] | 1
15 | [5 empty] | 1
16 | [5 empty] | 1
17 | [5 empty] | 1
18 | [5 empty] | 1
19 | [5 empty] | 1
20 | [5 empty] | 1
Sum
Results
Contributors used
Coverage mode (pick from the list) | [5 empty] | Fixed k
Coverage factor k to use with Fixed k | [5 empty] | 2
Effective degrees of freedom v_eff (Welch-Satterthwaite)
Coverage factor k used
Method: GUM (JCGM 100:2008). u_i = quoted value / divisor × ¦c¦. The contributors are combined by root-sum-square as uncorrelated inputs.
The sheet does the arithmetic. Choosing the contributors, distributions, sensitivity coefficients and degrees of freedom is your judgement.
Free template by Axiospec, calibration management software. axiospec.com/tools/uncertainty-budget-template

## Example (filled)
Example budget: 0-25 mm outside micrometer
ILLUSTRATIVE ONLY. These figures show how the sheet works. They are not a budget for any real micrometer.
Instrument / measurement |  | 0-25 mm outside micrometer
Measurand |  | Indication error at 25 mm, checked against a 25 mm gauge block
Units |  | µm |  |  |  |  | Units of the measurand. Each u_i is in these units once c is applied.
Prepared by
Date
# | Contributor (source) | Type (A/B) | Quoted value | Distribution | Divisor | Sensitivity coefficient c | Standard uncertainty u_i | Degrees of freedom v_i (blank = infinite) | u_i² | % contribution | u_i⁴ / v_i | Check
1 | Reference standard: 25 mm gauge block, certificate U = 0.10 µm at k=2 | B | 0.1 | Normal k=2 |  | 1
2 | Micrometer resolution 1 µm, entered as a 0.5 µm half-width | B | 0.5 | Rectangular |  | 1
3 | Repeatability: s = 0.8 µm from 10 readings | A | 0.8 | Normal k=1 |  | 1 |  | 9
4 | Temperature difference ±1 °C between micrometer and gauge block | B | 1 | Rectangular |  | 0.2875
5 | Anvil parallelism 0.3 µm | B | 0.3 | Rectangular |  | 1
6 | [5 empty] | 1
7 | [5 empty] | 1
8 | [5 empty] | 1
9 | [5 empty] | 1
10 | [5 empty] | 1
11 | [5 empty] | 1
12 | [5 empty] | 1
13 | [5 empty] | 1
14 | [5 empty] | 1
15 | [5 empty] | 1
16 | [5 empty] | 1
17 | [5 empty] | 1
18 | [5 empty] | 1
19 | [5 empty] | 1
20 | [5 empty] | 1
Sum
Results
Contributors used
Coverage mode (pick from the list) | [5 empty] | From effective degrees of freedom (95%)
Coverage factor k to use with Fixed k | [5 empty] | 2
Effective degrees of freedom v_eff (Welch-Satterthwaite)
Coverage factor k used
Row 4: the quoted value is in °C. c = 25 mm × 11.5 × 10^-6 /°C = 0.2875 µm/°C converts it to µm.
Row 3 is Type A: s from 10 readings, so 9 degrees of freedom. The other rows are left blank, which counts as infinite.
The coverage mode is set to the effective degrees of freedom to show how k follows from v_eff. Switch it to Fixed k to use k = 2.
Method: GUM (JCGM 100:2008). u_i = quoted value / divisor × ¦c¦. The contributors are combined by root-sum-square as uncorrelated inputs.
The sheet does the arithmetic. Choosing the contributors, distributions, sensitivity coefficients and degrees of freedom is your judgement.
Free template by Axiospec, calibration management software. axiospec.com/tools/uncertainty-budget-template